## Slide 1: FlexPatch Design Notes
5/29/2021

## Slide 2: Design requirements: Forearm (or lower leg)
Number of electrodes (channels): 32
  Arrangement: 4x8 grid
  Interelectrode spacing: 10-30 mm
  Ribbon cable length
Coverage area: ~150 mm x ~150 mm
  ~60% of distance from elbow to wrist x width of arm at widest point
  electrode footprint should be tapered from elbow to wrist
Anchoring
  Flex substrate needs to be anchored to the arm; start with elastic straps
  Multiple patches will be tiled together to cover the forearm or lower leg

## Slide 3: Concept 1
Elastic straps
EMG patch
4 x 32-channel patches
Rectangular grid (e.g. 4x80); no taper

## Slide 4: Concept 2
64-channel patch on anterior (palmar) side and 32-channel patch on posterior side
32-channel patch
  4x8 grid
64-channel patch
  will mate with MB via 2x32-channel ZIF connectors
  Tapered to accommodate extra width near elbow
Elastic straps
EMG patch

## Slide 5: Anchoring straps
.75” or 1” elastic straps (Amazon)
Attachment to flex substrate (options)
  Preferred: boards built with slits cut into edges; stiffeners will need to be added to the edges
  Alternative: screws or rivets attached to holes in flex
  Other ideas?
Straps should provide tension to hold the electrodes against the skin

## Slide 6: Patch Design Changes
Increase electrode spacing to 20mm
Add features for anchoring straps
Add ribbon cable to extend edge connector ~10 cm (make some measurements on your own arm to estimate distance)
NOTE: we will need 4 different patch designs as shown at right